SCINTREX

Head Office

SCINTREX Limited

222 Snidercroft Road

Concord, Ontario, Canada L4K 1B5

Tel: +1-905-669-2280

Fax: +1-905-669-6403

e-mail:scintrex@scintrexltd.com

website: www.scintrexltd.com

In the U.S.A.

SCINTREX Inc.

900 Woodrow Lane, Suite 100

Denton, Texas

76205 U.S.A.

Tel: +1-940-591-7755

Fax: +1-940-591-1968

e-mail:richardj@scintrexusa.com

In S.E. Asia

SCINTREX / AUSLOG

P.O. Box 125 Sumner Park

83 Jijaws Street. Brisbane, QLD

Tel: +61-7-3376-5188

Fax: +61-7-3376-6626

e-mail:auslog@auslog.com.au

website: www.auslog.com.au

Technical Information

PLEASE PASS THIS IMPORTANT NOTE TO ALL OPERATORS OF YOUR

CG-3 GRAVIMETERS

TECHNICAL NOTE #4 TO ALL CG-3 OWNERS

Re: Spring Service Life and Reading Range

The uncorrected gravity reading is obtained from equation (A.4). The gravity reading range has high and low end, the range of which is guaranteed to be 7000 mGal, but is normally much larger. The low end of the gravity measurement range of the CG-3 may be calculated by inserting value (SF/SC)=0.1 into the equation (A.4)

SUmin=0.1*GCAL1 + 0.01*GCAL2

The average value of SUmin is about 600 mGal. Depending on the values GCAL1 and GCAL2 this value may differ somewhat for your meter. It is recommended that you calculate this value for your meter, because the gravity readings of the lower values than SUmin, may not be accurate.

The high end of the gravity measurement range may be calculated by inserting value (SF/SC)=1.457 into the equation (A.4)

Sumax=1.457*GCAL1 + 2.123*GCAL2

The average value of SU. about is 8800 mGal. Depending on the values GCAL1 and GCAL2 this value may differ somewhat for your meter. It is recommended that you calculate this value for your meter, because the gravity readings of the higher value than Smax may not be accurate.

The CG-3 sensor spring position is adjusted during the sensor manufacture in such a way that, on the average, the gravity reading in Toronto (latitude 43.7 degrees) is about 4200 mGal. This, in conjunction with the gravity reading range of about 8000 mGals, assures that newly produced CG-3 may measure the relative gravity anywhere on the surface of the earth, from elevations of a few thousand meters above the equator to sea levels at both poles. For you convenience, values of normal gravity at sea level, as a function of latitude, are listed in Table I.

Because of the positive drift rate of the CG-3 gravity sensor, the uncorrected gravity readings at the fixed locations gradually increase in value. The linear drift of the CG-3 sensor is caused by an unavoidable creep of the quartz spring, whose length under tension increases, on the average, by 0.5 ppm/day at the operating temperature of 60°C. A typical sensor with the drift rate of 0.5 mGal/day will reach the high end of the measurement range, at the latitude of 60 degrees, in about 20 years. If used at lower latitude, as most gravimeters are, the typical CG-3 service life will be even longer.

You may estimate the remaining service life of your meter in the following way:

a)	Calculate the value SUmax for your meter.

b)	Obtain the gravity reading at your location.

c)	If you are planning to use your meter at higher latitudes than your present latitude, evaluate from Table I the gravity difference between your latitude and the highest latitude at which you will be using your meter. Add this difference to the reading obtained in step b) to obtain the value SUm.

d)	The remaining service life expressed in days is calculated by dividing the difference (SUmax-SUmin.) by the value of the linear drift rate for your CG-3.

E.g. you have calculated for your meter from equ. (A.4) SUmax=8800 mGal, and the drift rate of your meter is 0.51 mGal/day. Assuming that, at your latitude of 50 degrees, you have measured SU=5012 mGal, and that you will be using your meter up to latitudes of 60 degrees, so that SUm=(5012+847)mGal, then the number of remaining service days is (8800-5012-847)/0.51=5767 days.

If, after many years of service, your CG-3 has drifted out of range, then the spring position may be reset by a routine procedure at the Scintrex plant. You may use also this occasion to inquire about the available programs to overhaul your meter, to upgrade to latest firmware version or to upgrade to CG-3M accuracy level in case the evaluation at Scintrex proves that the quality of your sensor warrants such an upgrade.

45 46 47 48 49 50 51 52 53 54 55 56 57 58 59 60 61 62 63 64 65 66 67 68 69 70 71 72 73 74 75 76 77 78 79 80 81 82 83 84 85 86 87 88 89 90

1

2

3

4

5

6

7

8

9

10

11

12

13

14

15

16

17

18

19

20

21

22

23

24

25

26

27

28

29

30

31

32

33

34

35

36

37

38339

40

41

42

43

44

45

NORMAL GRAVITY AT SEA LEVEL (TORGE, 1989, p5l)

g=ge*(I+b*sin2(I)-bl*sin2(2* I), gravity at latitude I where: ge=978.03268 [GAL], gravity at equator b=0.0053-244

bl=0.0000058

I =latitude

LATITUDE

g (GAL)

g-ge(mGAL)

LATITUDE

Table I

g(GAL)

g-ge(mGAL)

978.032680 978.034253 978.038969 978.046823 978.057805 978.071902 978.089098 978.109371 978.132697 978.159048 978.188392 978.220695 978.255916 978.294015 978.224944 978.378655 978.425095 978.474208 978.525935 978.580214 978.636978 978.696160 978.757689 978.821489 978.887485 978.955596 979.025740 979.097832 979.171785 979.247510 979.324915 979.403907 979.484390 979.566266 979.649436 979.733799 979.819254 979.905696 979.993021 980.081122 980.169892 980.259224 980.349009 980.439138 980.529500 980.619987

0.000 1.573 6.289 14.143 25.125 39.222 56.418 76.691 100.017 126.368 155.712 188.015 223.236 261.335 302.264 345.975 392.415 441.528 493.255 547.534 604.298 663.480 725.009 788.908 854.805 922.916 933.060 1065.152 1139.105 1214.830 1292.235 1371.227 1451.710 1533.586 1616.756 1701.119 1786.574 1873.016 1960.341 2048.442 2137.212 2226.544 2316.329 2406.458 2496.820 2587.307

980.619987 980.710488 980.800892 980.891089 980.980970 981.070424 981.159343 981.247618 981.335140 981.421804 981.507502 981.592131 981.675586 981.757765 981.838568 981.917895 981.995650 982.071737 982.146062 982.218535 982.289066 982.357569 982.423960 982.488156 982.550080 982.609654 982.666806 982.721465 982.773564 982.823038 982.869828 982.913875 982.955125 982.993526 983.029033 983.061600 983.091188 983.117761 983.141285 983.161732 983.179075 983.193295 983.204373 983.212296 983.217053 983.218640

2587.307 2677.808 2768.212 2858.409 2948.290 3037.744 3126.663 3214.938 3302.460 3389.124 3474.822 3359.451 3642-906 3725.085 3805.888 3885.215 3962.970 4039.057 4113.382 4185.855 4256.386 4324.889 4391.280 4455.476 4517.400 4576.974 4634.126 4688.785 4740.884 4790.358 4837.148 4881.195 4922.445 4960.846 4996.353 5028.920 5058.081 5085.081 5108.605 5129.052 5146.395 5160.615 5171.693 5179.616 5184.373 5185.960